INFN sez. Padova. ICARUS group.          The signal interconnections…. of the  T600 detector.          13/03/02

[image: image1.wmf]CC-1-6

Power

supply

9

1

16

1

FANOUT

out

In

CCS-1-9

Crate NIM

9

1

out

In

CCS-1-9

9

1

out

In

CCS-1-9

9

1

out

In

CCS-1-9

16

1

FANOUT

16

1

FANOUT

16

1

FANOUT


The signal interconnections between the readout racks on the top of T600 detector.

Author: 

Baibussinov Bagdat 

Padova, 2002.

The last version of readout electronics is structured as four rows of rack on the top of detector.  Each rack hosts all the readout channels needed to handle the set of wires connected though the corresponding feed-through. For T600 detector we have the rack with 576 channels on each (Fig 1). . 

Figure 1. The readout rack of ICARUS detector












[image: image2.jpg]i

i
fidll





 Each  rack has: 

One Analogue Crate: 

- distribution and decoupling of the bias voltage to the wires (A764);

- test pulse distribution and setting of baselines (V793);

 - accommodation 18 analogue boards providing amplification, shaping and analogue to digital conversion (V791C or Q).

One Digital Crate:

- 18 digital boards, where waveforms are buffered and processed to archive zero skipping (V789).

- Distribution of absolute clock and trigger signals (V816);

- 1 embedded CPU, controlling data collection and forwarding data to event builder (VME2100).

One linear power supplies for Analogue Crate. 

One rack monitor, which monitors internal temperatures and ventilation flow, and allows remote control (through the I2C bus) of the power suppliers and line voltage.

The row of the racks has 18 normal racks and 4 modified racks. Two modified racks have just digital crates and another two have analogue crates with power suppliers.  

To reduce the length of the cables each row of the rack and to divide them electrically one crate NIM for each row is used. Inside of the crate several Fan-Out modules are accommodated.  Total number of crates NIM to distribute the common signal for ICARUS detector is four and they on the middle position of the rows are installed.  Below you find common description of this system and some suggestion to avoid the common problem during installation of it.

The rack has connections with:
 
-
Power rack  (line cable);









-
Monitor PC (I2C and CAENet interfaces);









-
Ethernet Switch (Ethernet cable);

-
Crate NIM (Test pulse, Trigger, Global Clock, Enable signals).

Crate NIM.
The crate NIM use to distribute all commons signals to analogue e digital crates of one row of rack.  This crate receive a few cables from trigger system and control room:

1. Trigger signal.

2. Enable 

3. Reset.

4. Test pulse

and more number of cables have been distributing from crate NIM:  














Number of Modules



a. 24-25 cables of the Test pulse









from modules  CCS-1-9










3

b. 1-2 cables of the Clock 40 MHz.



 from module CC-1-6










1

c. 24-25 cables of the Global Clock 20 MHz.







1 

from module FANOUT

d. 24-25 cables of the RESET CPU









1 

from module FANOUT

e. 24-25 cables of the TRIGGER










1 

from module FANOUT

Compact version of NIM crates with 10 slots is used. Crate NIM has 4 modules CCS-1-9, one module CC-1-6, 4 modules DUAL FANOUT (1-16) and the power supply too (Fig.2). 

Total power of the crate:

Quantity  
 Current 


  

Power                                                                                            +6 Volt          -6 Volt

1. Modules  CCS-1-9                          4                       125 mA           125 mA           < 5 W 

2. Modules CC-1-6                             1                        130 mA           160 mA            < 5 W 

3. Modules FANOUT 1-16                 4                          

[image: image3.wmf]V1

K1

X3

R1

97

X1

R2

97

X4

C1

1pF

C2

1pF

X5

R3

50

R4

50

T1

W0

W1

W2

W3

X6

X7

X8

Shielded tweasted pair

4

3

2

1

From 2,3,4 modules CCS1-9 to another racks (max number 27 racks)

input of V791C(Q)

input of A764

The crate NIM.

The readout rack

V793

Test pulse generator

Ccal =1 pF

Ccal =1pF

to A764

to V791

CCS1-9

Out9

Out1

Out9

Out1

Out9

Out1

Out2

Out3

Out9

Out1

CCS1-9

CCS1-9

CCS1-9

Out2

Out4

Shielded tweasted pair

 
Figure 2. The layout of the crate NIM.  

Distribution of the test pulse.
There are four kinds of calibration that by V793 modules are provided:

(See the V793 specification).




Input of


note
 

Calibration 1

V791C, Q
  


test just V791

Calibration 2 

A764 

  


test A764, V793 together

Calibration 3 

test point 1 



begin of wires, used to test 

connection of wires 

Calibration 4 

test point 2



on the end of wires, used to test  

channel with cable and all boards.

The module CCS 1-9 receive signal from a pulse generator (Fig.3). Normally the module allows transmit any waveform signal and we can use a custom waveform generator. The choice of kind of signal depends from test conditions.  To verify the coefficients of the channels and signal to noise ratio we can use step pulses for different calibration test.  

All calibrations have different coefficients because of the noise and interference between different signals on the backplane of the analogue crate was found. 

The table 1 shows all calibration coefficients from V793 to test points, which must be used during calculation of channels characteristic.   

Table 1.


Gain V793
Gain  V791
Gain A764
on the test capacitance(Count/fC)




V791C
2.3

0.1


Count/(0.23*Vin(mV))*1pF


A764

2.3


2
Count/(4.6*Vin(mV))*1pF


Test point1
2.3

0.1
Count/(0.23*Vin(mV))*1pF


Test point 2
25/22=0.114


Count/(0.114*Vin(mV))*1pF


Note. Attention please the signal on the input V793 is differential. 

Obviously, the response of the amplifier with shaper will depends from waveform test signal. If you have the test pulse like step voltage with fast rise time we find one coefficient. And for triangle signal the coefficient will be different. 

We calculate the gain of the analogue board for step signal  5+/-0.5 count/fC  for V791C and 6.5+/- 0.5 counts/fC for V791Q boards. 

[image: image4.wmf]CC-1-6

Power

supply

9

1

16

1

FANOUT

out

In

CCS-1-9

Crate NIM

9

1

out

In

CCS-1-9

9

1

out

In

CCS-1-9

9

1

out

In

CCS-1-9

16

1

FANOUT

16

1

FANOUT

16

1

FANOUT


Figure 3. The scheme of the distribution of the test pulses.

All Cables.










FROM 


TO


1. Power cable  







Power rack            Rack

2. Interface I2C







PC 


Monitor Module

or 

Monitor Module 

Monitor Module

3. 
CAENet interface cable               


PC 



V793

or   

             V793

        
V793

3. Clock  20 MHz.


  



FANOUT

V816

4. RESET








FANOUT 

V816


5.  Trigger Signal







FANOUT 

V816

6.  Enable 








FANOUT 

V816

7.  Calibration Pulse 

 




CCS-1-9 

V793

8. Clock 40 MHz.






CC-1-6


V793

9.  Ground cable 







Flange 

   
Rack

10. High Voltage 1




Power Supply 1


HVfilter

11. High Voltage 2




Power Supply 2


Hvfilter

Inside of rack:

12. 18 signal cables 






Flange 


V764

13.  4 test signals


 




Flange 


V793

14. 18 cables  of link






V791


V789

All boards:

Module:




Function





Number total

V789




Digital board.





18

V791
 



Analogue board.




18


V793


 

Slow control Module. 



1

V764




Decoupling Board. 



18

V816
 



Trigger Module. 



 
1

Hvfilter




Filter of High Voltage



1

Connection board










1

Specification of the signals.
1. RESET, Global Clock 20 MHz, Trigger, Enable signals are standard NIM.

2. Clock 40 MHz is LVDS standard signal. To transfer the signal twisted pair is used.


3. Calibration signal is differential analogue signal. The maximum of signal is +/- 3 Volts.

4. CAENet interface signal is isolated NIM signal.

� EMBED Word.Picture.8  ���















2

_1074520742.doc


CC-1-6







Power 







supply







9







1







16







1







FANOUT







out







In







CCS-1-9







Crate NIM







9







1







out







In







CCS-1-9







9







1







out







In







CCS-1-9







9







1







out







In







CCS-1-9







16







1







FANOUT







16







1







FANOUT







16







1







FANOUT












